## Slide 1: Bubble CPAP in the NICU
Fu-Sheng Chou | Michael Doti
Kaiser Permanente Riverside Medical Center

## Slide 2: Nasal continuous positive airway pressure (NCPAP)
Noninvasive
Maintains lung volume (functional residual capacity, FRC)
Reduce the need for intubation
Reducing morbidities associated with mechanical ventilation

## Slide 3: Pressure = Flow x Resistance

## Slide 4: Ventilator-derived NCPAP (V-NCPAP)
Use both flow and resistance (valve) to generate desired pressure
CPAP level is maintained at the distal expiratory-limb orifice
Adjusts to keep the delivered CPAP at the desired level

## Slide 5: Bubble NCPAP (B-NCPAP)
Introduced in the 1970’s
Flow is also constant
Expiratory limb of the CPAP tubing is immersed in an underwater chamber
Depth in centimeters equal to the ‘desired’ cmH2O CPAP level
Flow of gas through the system creates bubbling in the chamber

## Slide 6: Bubble NCPAP (B-NCPAP)
Favored by an institution reporting a low incidence of BPD
Improved gas exchange and lung recruitment during B-NCPAP, due to the high-frequency oscillatory content of the bubbling
More flow → more bubbling → better oscillation → better gas exchange
Lower FiO2, lower length of stay, less invasive ventilation
(demonstration)

## Slide 7: Bubble NCPAP (B-NCPAP)
Gas needs to get out, pressure built up to the depth of the tubing to overcome water pressure
Oscillatory effect from bubbling

## Slide 8: How does bubble CPAP work?

## Slide 9: Fisher & Paykel bubble CPAP (bCPAP) system

## Slide 10: Work of breathing variance

## Slide 11: CO2 retention variance

## Slide 12: Outcomes
Success is defined as weaning off CPAP due to clinical improvement
Failure: FiO2 > 60%, CPAP ≥ 8 cmH2O, apnea requiring PPV, shock, intubation

## Slide 13: Factors affecting the success of B-NCPAP
Nasal interface
  Nasal prong, nasal mask (closed, low resistance)
  RAM Cannula (open, high-resistance; if no leak, significantly increased WOB)

## Slide 14: Factors affecting the success of B-NCPAP
Monitoring and preventing nasal injury
  Frequent assessment
  Risk of losing FRC during assessment
Fixation of the nasal interface
  Avoid lateral pressure, especially on the nasal septum

## Slide 15: Nasal injury scoring chart

## Slide 16: Factors affecting the success of B-NCPAP
Bubbling and system check
  No flow, no bubbling
  Tubing too deep because water builds up
Optimizing pressure delivery
  Pacifier / chin strap
  OG tube to vent

## Slide 17: Compared to V-NCPAP, B-NCPAP does not…
Reduce mortality
Reduce pneumothorax
Reduce adverse neurodevelopmental outcomes
Reduce BPD

## Slide 18: Q&A